COB1474	AN AUGMENTED ENDURANCE TEST WITH ELECTROMECHANICAL PULSATILE TOTAL ARTIFICIAL HEART (TAH)

Yukio Ohashi, Aron de Andrade, Kin-ichi Nakata, Jürgen Müller, Yukihiko Nosé.

Baylor College of Medicine, Department of Surgery, One Baylor Plaza, Houston, Texas 77030, USA - E-mail: yohashi@bcm.tmc.edu

A Total artificial heart (TAH) has been developed in the Department of Surgery at Baylor College of Medicine. This TAH successfully underwent a hydrodynamic study, hemolysis study and two acute implantations. Before initiating a preclinical study, the reliability of the device needs to be documented based on the durability of each component including the electrical parts through a long-term endurance study,

This TAH was submitted to a four-month running test with the most severe driving condition where a maximum dP/dt was chosen more than 15,000 mm Hg/sec that is almost six times higher than that of a normal driving condition. The pump was submersed into a water bath that was maintained at 37O C. The TAH was driven in a LMA variable rate mode at 8 L/min outflow, 15 mm Hg preload, 100 mm Hg left afterload and 25 mm Hg right afterload. The outflow, pressures, temperature inside the TAH and electrical current were monitored.

Several stress concentrating portions were detected. The connection between the roller nut and support plate lies in one of the most stress concentrating portions, and a more reliable fixing procedure was required, A redesign was made in this portion to offer a durable connection. No malfunction of the actuator and controller was detected throughout the testing duration. The required electrical power was 14.7 ± 1.6 watts and the total efficiency was 12.6 ± 1,4 %. The temperature in the TAH was 45.9 ± 0.9O  C on the motor, and 37.8 ± 0.3O C on the centerpiece.

Keywords: total artificial heart, durability, augmented endurance study, stress concentrating portion.