cob853	Avaliação da transferência de oxigênio em oxigenadores de sangue utilizando água no lugar do sangue / Oxygen transfer evaluation of blood oxygenators with water

Fábio Turri & Adolfo A. Leirner

Divisão de Bioengenharia - InCor – HCFMUSP – Av. Dr. Enéas de C. Aguiar, 44 – CEP 05403-000 São Paulo – SP, Brasil

E-mail: biofabio@incor4.incor.usp.br

In vitro assays with bovine blood are a current practice in blood oxygenators (artificial lung) development. The aim of this work is to show that it is possible to use water in those tests instead of blood and obtain a correlation that can be used to predict blood-gas oxygen mass transfer rate. Water tests have the advantage to avoid logistic problems regarding blood supply. Four commercial blood oxygenators were firstly tested with water at the following condition: flow = 6 l/min., temperature = 37 oC,    and inlet pO2 equal to 30 mmHg. Oxygen 21 %    was vented in a 1:1 blood-gas flow relationship. Each oxygenator was secondly tested with bovine fresh blood at    a flow of 6 l/min., and 100 % oxygen at 1:1 blood-gas flow relationship. Blood parameters were set according to the AAMI conditions. Oxygen mass transfer effectiveness coefficients (Ef) were calculated.    Statistical analysis shows linear correlation with r2= 0,927 and Efblood = 1,95*Efwater-1,02.

Keywords: Blood oxygenation, membrane oxygenation, cardiopulmonary bypass/ Oxigenadores de sangue, oxigenadores de membrana, circulação extra-corpórea.