COB1157	TECHNIQUE FOR ASSESSING THE PERFORMANCE OF CO-ORDINATE MEASURING MACHINES (CMMs)

Joao  Bosco de Aquino Silva

Departamento de Engenharia Mecanica, Universidade Federal da Paraiba

CEP 58100-000  Campina Grande, Pb, Brasil - E-mail: jboscona@terra.npd.ufpb.br

M. Burdekin

Department of Mechanical Engineering, University of Manchester Institute of Science and Technology, UMIST, UK

A new technique is proposed for assessing the performance of co-ordinate measuring machines(CMMs). This technique comprises a new form of space frame which has the form of a tetrahedron and is referred to as modular space frame as different configurations of tetrahedron can be obtained easily and rapidly. The modular space frame comprises magnetic ball links and a ball plate. The volumetric error data obtained when a CMM measures  the calibrated modular space frame can be used to verify whether a CMM maintains the manufacturer specifications. The experimental results have demonstrated that the modular space frame technique has an acceptable repeatability and provides a practical and cost effective mechanical artefact for calibration, verification, periodic reverification and  acceptance test of any type of CMM.

Keywords: Calibration; Verification; Co-ordinate Measuring Machines